## Multiplace
|  |  |  | Document #: |  | qa document no.
DEPTH GAUGE CALIBRATION: |  |  |  |  | Facility, location
Performed on: |  | date | Using: | Model |  | Digital Master Gauge: |  |  | Serial number
Performed by: |  | technician's name |  |  |  | Gauge Full Scale: FSW |  |  | 230
Master |  |  | Main lock |  |  |  |  | Entry Lock |  |  |  |  | Therapeutic Gauge
psi | FSW | Target | FSW | Actual deviation |  | Allowed |  | FSW | Actual deviation |  | Allowed |  | FSW | Actual deviation |  | Allowed
0 | 0 | 0 | 0 | 0 | 0% | ± | 0.25% | 0 | 0 | 0% | ± | 0.25% | 0 | 0 | 0% | ± | 0.25%
7.25 | 16.2811 | 15 | 16 | 0.28109 | 0.1222% | ± | 0.25% | 16 | 0.28109 | 0.1222% | ± | 0.25% | 16.3 | -0.018912 | -0.008223% | ± | 0.25%
11.75 | 26.3866 | 25 | 26 | 0.38659 | 0.1681% | ± | 0.25% | 26 | 0.38659 | 0.1681% | ± | 0.25% | 26.4 | -0.01341 | -0.00583% | ± | 0.25%
20.5 | 46.0362 | 45 | 46 | 0.036179 | 0.01573% | ± | 0.25% | 46 | 0.036179 | 0.01573% | ± | 0.25% | 46 | 0.036179 | 0.01573% | ± | 0.25%
26.25 | 58.9488 | 60 | 59 | -0.051234 | -0.02228% | ± | 0.25% | 58 | 0.94877 | 0.4125% | ± | 0.25% | 59 | -0.051234 | -0.02228% | ± | 0.25%
29.25 | 65.6858 | 65 | 66 | -0.31423 | -0.1366% | ± | 0.25% | 66 | -0.31423 | -0.1366% | ± | 0.25% | 66.7 | -1.0142 | -0.441% | ± | 0.25%
43.75 | 98.2479 | 100 | 98 | 0.24794 | 0.1078% | ± | 0.25% | 99 | -0.75206 | -0.327% | ± | 0.25% | - | - | - | - | -
72.75 | 163.3723 | 165 | 165 | -1.6277 | -0.7077% | ± | 0.25% | 164 | -0.62771 | -0.2729% | ± | 0.25% | - | - | - | - | -
43.75 | 98.2479 | 100 | 98 | 0.24794 | 0.1078% | ± | 0.25% | 99 | -0.75206 | -0.327% | ± | 0.25% | - | - | - | - | -
29.25 | 65.6858 | 65 | 66 | -0.31423 | -0.1366% | ± | 0.25% | 65 | 0.68577 | 0.2982% | ± | 0.25% | 65.7 | -0.014232 | -0.006188% | ± | 0.25%
26.25 | 58.9488 | 60 | 59 | -0.051234 | -0.02228% | ± | 0.25% | 58 | 0.94877 | 0.4125% | ± | 0.25% | 59 | -0.051234 | -0.02228% | ± | 0.25%
20.5 | 46.0362 | 45 | 46 | 0.036179 | 0.01573% | ± | 0.25% | 45 | 1.0362 | 0.4505% | ± | 0.25% | 46 | 0.036179 | 0.01573% | ± | 0.25%
11.75 | 26.3866 | 25 | 26 | 0.38659 | 0.1681% | ± | 0.25% | 56 | -29.6134 | -12.88% | ± | 0.25% | 26.4 | -0.01341 | -0.00583% | ± | 0.25%
7.25 | 16.2811 | 15 | 16 | 0.28109 | 0.1222% | ± | 0.25% | 16 | 0.28109 | 0.1222% | ± | 0.25% | 16.3 | -0.018912 | -0.008223% | ± | 0.25%
0 | 0 | 0 | 0 | 0 | 0% | ± | 0.25% | 0 | 0 | 0% | ± | 0.25% | 0 | 0 | 0% | ± | 0.25%
Notes: |  | Master Gauge calibration certificate required to be appended to this report.
Report on significant deviations as these apply to the treatment tables being used.
Commercial diving accuracies are not relevant to air/oxygen treatment table depth readings.
Deviations should be relevant to what can be reliably read by eye on the analogue gauges.

## Monoplace
Document #: |  | qa document no.
PRESSURE GAUGE VALIDATION: | [5 empty] | Facility, location
Performed on: |  | date |  | Using: | Model |  | Master Gauge
Gauge Full Scale: |  | 30 | psi |  | Gauge serial #: |  | Serial number
Performed by: |  | technician's name |  |  | Gauge calibr date: |  | Gauge calibration date
Master |  |  |  | Monoplace Chamber
psi | ATA | Target |  | psi | ATA | Actual deviation |  | Allowed
0 | 0 | 0 |  | 0 | 1 | psi | 0% | ± | 0.25%
4.41 | 1.3001 | 1.3 |  | 4.4 | 1.2994 | psi | 0.03333% | ± | 0.25%
7.35 | 1.5001 | 1.5 |  | 7.3 | 1.4967 | psi | 0.1667% | ± | 0.25%
11.77 | 1.8009 | 1.8 |  | 11.8 | 1.8029 | psi | -0.1% | ± | 0.25%
14.7 | 2.0003 | 2 |  | 14.7 | 2.0003 | psi | 0% | ± | 0.25%
20.57 | 2.3997 | 2.4 |  | 20.6 | 2.4017 | psi | -0.1% | ± | 0.25%
26.45 | 2.7998 | 2.8 |  | 26.5 | 2.8032 | psi | -0.1667% | ± | 0.25%
29.39 | 2.9999 | 3 |  | 29.4 | 3.0006 | psi | -0.03333% | ± | 0.25%
35.27 | 3.4 | 3.4 |  | 35 | 3.3816 | psi | 0.9% | ± | 0.25%
29.3 | 2.9937 | 3 |  | 29.4 | 3.0006 | psi | -0.3333% | ± | 0.25%
26.5 | 2.8032 | 2.8 |  | 26.5 | 2.8032 | psi | 0% | ± | 0.25%
20.5 | 2.3949 | 2.4 |  | 20.5 | 2.3949 | psi | 0% | ± | 0.25%
14.7 | 2.0003 | 2 |  | 14.7 | 2.0003 | psi | 0% | ± | 0.25%
11.8 | 1.8029 | 1.8 |  | 11.8 | 1.8029 | psi | 0% | ± | 0.25%
7.3 | 1.4967 | 1.5 |  | 7.3 | 1.4967 | psi | 0% | ± | 0.25%
4.4 | 1.2994 | 1.3 |  | 4.4 | 1.2994 | psi | 0% | ± | 0.25%
0 | 1 | 0 |  | 0 | 1 | psi | 0% | ± | 0.25%
Notes:
Master Gauge calibration certificate required to be appended to this report.
Report on significant deviations as these apply to the treatment tables being used.
Commercial diving accuracies are not relevant to air/oxygen treatment table depth readings.
Deviations should be relevant to what can be reliably read by eye on the analogue gauges.